
[image: image36]
[image: image37]

CONTENT

- 1 -I. Safety rules and precautions


- 1 -II．Introduction


- 1 -III. Range and accuracy


- 2 -IV．Model of Series


- 2 -V．Technical Specification


- 3 -VI．Structure


- 4 -VII．Operation


- 4 -1. Switch on/off


- 4 -2. Battery voltage check


- 4 -3. Resistance precision test


- 5 -4. Resistance precision test


- 6 -5. Backlight control


- 6 -6. Data lock/storage


- 6 -7. Data access/delete


- 7 -8. Data upload


- 7 -9. Line resistance calibration (removal of residual resistance)


- 8 -10. Battery description


- 8 -VIII. Accessories




I. Safety rules and precautions

Thank you for purchasing our company's DC low resistance meter. In order to better use this product, please be sure to:

——Read this user manual in detail.

——Strictly follow the safety rules and precautions listed in this manual.

 In any case, special attention should be paid to safety when using this tester.

Can’t to measure live objects. Before measuring, make sure that the measured resistance or metal body is not charged, otherwise there is a danger of burning the instrument.

The battery low voltage symbol shows, please charge it in time, 5-8 hours per charge.

If the meter is left unused for a long time, please charge the battery every 1-2 months.

Please stop using the meter when the test wire is broken and the metal is exposed.

Do not place or store the tester for a long time in a place with high temperature and humidity, condensation and direct sunlight.

Precision instruments must be maintained regularly. Keep the body and test lines clean and do not drop them.

Use, dis assembly and maintenance of this tester must be operated by authorized personnel.

Due to the reason of this tester, if it is dangerous to continue using it, it should be stopped immediately, sealed up immediately, and handled by an authorized organization.

[image: image38.emf]The user must follow the instructions for safe operation of the "  " hazard signs on the tester and manual.
II．Introduction
FR3060 Digital DC low resistance meter, also known as micro-ohmmeter, ohmmeter, equipotential tester, using the microprocessor technology, four-wire test, safety, precision and reliability.Mainly used to measure the wire resistance of cables, the contact resistance of switches, connectors, relays, the resistance of coils, motors, transformer windings, and equipment enclosures, lightning protection belts, ground beams, structures, cabinets, steel bars, water pipes, windows, guardrails, and heat dissipation Connection resistance test between metal components of objects such as devices, assembly lines, etc.

It is widely used in telecommunications, electric power, meteorology, computer rooms, oil fields, power distribution lines, tower transmission lines, gas stations, factory grounding grids, lightning rods, etc.

FR3060 digital DC low resistance meter is composed of host, monitoring software, test line, communication line, etc. The large-screen LCD display of the main unit is clear at a glance, and the port overload prevention function prevents the instrument from being burned out by misoperation.Store 500 sets of data, resistance measurement range: 0.001mΩ~ 300.0KΩ. The upper computer software has functions such as historical data reading, consulting, saving, and reporting.

III. Range and accuracy

	Model
	Range
	Accuracy
	Resolution
	Maximum test current

	FR3060
	0.001mΩ～10.000mΩ
	Within 18℃~28℃, Below 70%rh:
±0.1%FS±20dgt
	0.001mΩ
	1.2A

	
	10.01mΩ～100.00mΩ
	
	0.01mΩ
	1.2A

	
	100.1mΩ～1000.0mΩ
	
	0.1Ω
	1.2A

	
	1.001Ω～10.000Ω
	
	0.001Ω
	1.2A

	
	10.01Ω～100.00Ω
	
	0.01Ω
	0.5A

	
	100.1Ω～1000.0Ω
	
	0.1Ω
	0.05A

	
	1.001KΩ～10.000KΩ
	
	0.001KΩ
	5mA

	
	10.01KΩ～100.00KΩ
	
	0.01KΩ
	0.5mA

	
	100.1KΩ～300.0KΩ
	
	0.1KΩ
	0.05mA

	FR3060A
	0.01mΩ～100.00mΩ
	Within 18℃~28℃, Below 70%rh:
±0.1%FS±20dgt
	0.01mΩ
	1.2A

	
	100.1mΩ～1000.0mΩ
	
	0.0001Ω
	1.2A

	
	1.001Ω～10.000Ω
	
	0.001Ω
	1.2A

	
	10.01Ω～100.00Ω
	
	0.01Ω
	0.5A

	
	100.1Ω～1000.0Ω
	
	0.0001KΩ
	0.05A

	
	1.001KΩ～10.000KΩ
	
	0.001KΩ
	5mA

	
	10.01KΩ～30.00KΩ
	
	0.01KΩ
	0.5mA

	FR3060C
	0.01mΩ～100.00mΩ
	Within 18℃~28℃, Below 70%rh:
±0.1%FS±20dgt
	0.01mΩ
	1.2A

	
	100.1mΩ～1000.0mΩ
	
	0.0001Ω
	1.2A

	
	1.001Ω～10.000Ω
	
	0.001Ω
	1.2A

	
	10.01Ω～100.00Ω
	
	0.01Ω
	0.5A

	
	100.1Ω～1000.0Ω
	
	0.0001KΩ
	0.05A

	
	1.001KΩ～10.000KΩ
	
	0.001KΩ
	5mA

	
	10.01KΩ～10.00KΩ
	
	0.01KΩ
	0.5mA


Note: within 18℃~28℃, above 70%rh, within -10℃~50℃, below 80%rh: ±0.2%FS±20dgt.
IV．Model of Series

	Model
	FR3060A
	FR3060
	FR3060C

	Resistance Range
	0.01mΩ～30.00KΩ
	0.001mΩ～300.0KΩ
	0.01mΩ～10.00KΩ

	Resolution
	10uΩ
	1uΩ
	10uΩ

	Accuracy
	±0.1%FS
	±0.1%FS
	±0.1%FS

	Data storage
	500 sets
	500 sets
	500 sets

	USB Data upload function
	N
	Y
	N

	Charging function
	Y
	Y
	Y


V．Technical Specification

	Function
	Mainly used to measure the wire resistance of cables, the contact resistance of switches, connectors, relays, the resistance of coils, motors, transformer winding and metal riveting resistance, the connection resistance test between metal components, the low-value resistance test, and the grounding grid Resistance test of connecting conductor, contact resistance test, etc.

	Detection method
	Four-wire method test

	Test current
	≥1A

	Open circuit voltage
	≤4.2V

	Overload protection
	Y

	Power
	Measuring power≤8W

	power supply
	DC 3.7V 2500mAh

	Backlight
	Controllable gray and white screen backlight, suitable for dim places

	Display Model 
	LCD

	LCD Size
	72 mm×55mm

	Instrument size
	187 mm×119 mm×50mm

	Test line length
	60cm，red, black 1pcs

	measure time
	About 2 times/sec

	USB interface
	With micro USB interface

	Communication line
	micro USB 1pcs

	Data storage
	500 groups, "MEM" storage indication, display "FULL" symbol to indicate storage is full

	Data review
	Data review function：“MR”symbol display

	Overflow display
	Over-range overflow function: "OL" symbol display

	Battery voltage
	Real-time display of battery power, reminding timely charging when battery voltage is low

	Automatic shut-down
	"APO" indicates, automatic shutdown after 15 minutes

	Power consumption
	Standby: about 100mA (backlight off)

	
	Turn on backlight: about 105mA

	
	Measure：2A Max(backlight off)

	Weight
	Instrument:480g(including battery)

	
	Test line：250g

	Working temperature and humidity
	-10℃～50℃；below 80%rh

	Storage temperature and humidity
	-20℃～60℃；below 70%rh

	Overload protection
	AC 220V/0.0001 second between C1-C2 P1-P2 ports, after overload protection, you need to restart the meter to test normally

	Insulation resistance
	More than 10MΩ(Between circuit and case 500V)

	Pressure resistance
	AC 3700V/rms(Between circuit and case)

	Electromagnetic properties
	IEC61010-4-3，Wireless frequency electromagnetic field≤1V/m 

	Suitable for Safety Regulations
	IEC61010-1, CAT III 600V, Pollution Level 2, JJG724-1991 "Determination of Direct Current Digital Ohmmeter", JJG166-1993 "Dialect Resistor Verification Regulation", "DL/T967-2005 Loop Resistance Tester and DC Resistance Fast Tester Verification Procedures


[image: image39.wmf] VI．Structure
1. USB/charging interface

charging indicator

Test line interface

LCD 

Rubber insulated protective sleeve

Function button area

Test line ( red and black each 2pcs)

VII．Operation

Switch on/off 
Press " 
[image: image1.emf] "  to switch the machine. APO is displayed in the lower corner after power on, and it will automatically power off after 15 minutes when it is not in operation.

 Battery voltage check
Wait for 2~4 seconds after powering on. If the LCD displays the low battery voltage symbol "[image: image2.wmf]", it means the battery is low, please charge it in time. Adequate battery power can ensure the accuracy of the measurement. When the battery power decreases, the power indicator bar decreases.

Resistance precision test

	[image: image3.wmf]
	During the test, first remove the insulating layer and oxide layer on the surface of the tested object.

	
	When measuring resistance or DC low resistance, live test is not allowed. Live test may burn the meter.

	
	The test clip may be oxidized after using it for a long time. Pay attention to clean the oxide or foreign matter attached to the clip to make the clip contact well.

	
	The connection between the test line, the detector and the object under test must be reliable.

	
	Since component test heating will cause errors, it is recommended that the test time should not exceed 30 seconds, and the interval between each test should be 30 seconds, the measurement results will be more accurate.

	
	The meter displays the OL symbol during the test, indicating that the resistance between the two points under test exceeds the range. It is necessary to restart the meter and test again to eliminate the fault caused by the over voltage protection. If the fault is caused by the over voltage protection, it means that the resistance being measured is charged, and the resistance to be measured should be cut off immediately, and the test should be performed after restarting the meter. Or check whether the test line is in good contact. It may be an open circuit between the two points being tested.


Resistance precision test
After power on, connect the wire as shown in the figure below to test the resistance value of the measured resistance.

[image: image4.wmf]
The following figure shows the example to 

test the resistance value between the meter box and the ground and the lower line.

[image: image5.wmf]
The following figure shows an example to test the resistance value of two metal connectors.

[image: image6.wmf]
The following figure shows an example to test the resistance value of the contact point of the relay.

[image: image7.wmf]
Backlight control

After booting, press the "
[image: image8.emf]" key to turn on or off the backlight. The backlight function is suitable for dim places. The default backlight is off every time you turn on
Data lock/storage

After turning on or after the measurement is completed, press the "
[image: image9.emf]" key to lock the currently displayed data, press the "
[image: image10.emf]" key to automatically number and store the current data, if the storage is full, the meter displays the "FULL" symbol. As shown in the figure below: the measured data is 70.11mΩ, press "
[image: image11.emf]" to display the latched data.


[image: image12.emf]  
[image: image13.emf]  
[image: image14.emf]
Data access/delete

After turning on or after the measurement, press the "
[image: image15.emf]" key to switch to the data access mode, and the word "MR" is displayed below. Press the “
[image: image16.emf]” or “
[image: image17.emf]” key to set the step value to the corresponding data of the storage data group of 1 group, and press the “
[image: image18.emf]” or “
[image: image19.emf]” key to set the step value of the corresponding data to the storage data group of 10 groups. Press "
[image: image20.emf]" or press "
[image: image21.emf]" twice to exit this mode and return to the test mode, as shown in the figure below.

When checking, the small number in the figure below is the number of currently checked groups. If there is no stored data, the LCD will display “NULL”, see the figure below.


[image: image22.emf]                             
[image: image23.emf]
In the data access mode, press the " 
[image: image24.emf] " key to switch to the data delete mode. , Press 

“
[image: image25.emf]” or “
[image: image26.emf]” to select “NO” or “YES”, select “NO” and then press “ 
[image: image27.emf] ” to return to the test state without deleting, select “YES” and press “ 
[image: image28.emf] ” to delete the stored data. After deleting, the following figure will be displayed .

[image: image29.emf]    



   
[image: image30.emf]
Data upload

Connect the USB communication line between the computer and the instrument, turn on the instrument, and run the host computer software. If the USB connection is successful, the stored historical data can be read, uploaded to the computer and saved.

The upper computer software has functions such as historical data reading, consulting, and saving.

Line resistance calibration (removal of residual resistance)

Short-circuit the two clips first, wait for the displayed value to stabilize, and finally press the "
[image: image31.emf]" button for 2 to 3 seconds to complete the line resistance calibration. As shown below:

Note: You must wait for the displayed value to stabilize before long pressing the "[image: image32.emf]" key to calibrate the line resistance.

[image: image33.wmf]
Battery description

	[image: image34.wmf]
	Generally charge for 5 to 8 hours.

	
	If the detector is left unused for a long time, please charge the battery every 1 to 2 months.

	
	Please use the original charger for charging. This machine does not support fast charging.

	
	When charging, the charger's red light is on, and the green light is on when it is fully charged.


The meter is powered by a 3.7V lithium battery. When the battery power is reduced, the power indicator bar will decrease and the power symbol “ [image: image35.wmf] ” will be displayed. Please charge it in time. The measurement accuracy is affected when the voltage is low.

VIII. Accessories
	Instrument
	1PC

	Instrument box
	1PC

	Monitoring Software CD
	1PC（only for FR3060）

	USB Communication line
	1PC

	Test line
	Kelvin test line 2 (red 60cm, black 60cm）

	Charger
	1PC

	Manual, certificate
	1SET


[image: image40.emf]

The contents of this user manual cannot be used as a reason to use the product for special purposes.

The company is not responsible for other losses caused by the use.

The company reserves the right to modify the contents of the user manual. Any changes will be made without further notice.

[image: image41.png]Fuzrr




GuangZhou ZhengNeng Electronics Technology CO.,LTD
Address: 4th Floor, No. 771, Guangcong Eighth Road, Changyaoling Village, Zhongluotan Town, Baiyun District, Guangzhou, China,Post code：510540
Tel：+86-020-36544172       

E-mail: sales@fuzrr.com﻿

Website：http://www.fuzrr.com
� EMBED CorelDraw.Graphic.20 \* MERGEFORMAT ���





� EMBED  \* MERGEFORMAT ���








PAGE  

_1234567897.unknown

_1234567905.unknown

_1234567909.unknown

_1234567913.unknown

_1234567917.unknown

_1234567919.unknown

_1234567920.unknown

_1234567921.unknown

_1234567918.unknown

_1234567915.unknown

_1234567916.unknown

_1234567914.unknown

_1234567911.unknown

_1234567912.unknown

_1234567910.unknown

_1234567907.unknown

_1234567908.unknown

_1234567906.unknown

_1234567901.unknown

_1234567903.unknown

_1234567904.unknown

_1234567902.unknown

_1234567899.unknown

_1234567900.unknown

_1234567898.unknown

_1234567893.unknown

_1234567895.unknown

_1234567896.unknown

_1234567894.unknown

_1234567891.unknown

_1234567892.unknown

_1234567890.unknown

